TYPICAL SPECIFICATION

HYDRAULIC CENTURY®

CARBOY SCALE  (200 - 600 lbs.)

SPECIFICATION SECTION 11245 - Chemical Feeders

SPECIFICATION SECTION 11240 - Chemical Metering Pumps:

Carboy scale shall be of the hydraulic load cell type.  Scale shall be sized to accept drum or carboy diameters from ______ to _______ inches (cm) (see chart below) .  Platform height shall be no more than 1-5/8" (38mm) to allow easy handling and unloading of drums.  Platform shall be constructed of 1" (25mm) PVC Plastic.   Platform scale coating system shall be a minimum dry thickness of 80 mils and be resistant to moisture, chemicals, abrasion, impact and UV light.

Load cell shall be of the temperature-stable, rolling diaphragm type.  A _____ ft. length of tubing (6'/2M standard) shall lead from the load cell to allow easy remote installation of a wall mounted dial.   Load cell system shall require no electric power and shall be immune to all RFI/EFI, power failures and lightening strikes.

Dial diameter shall be 4-1/2" (114mm) and read zero to_______ lbs. (kgs) with provision for tare adjustment.  Dial shall be temperature stable with damper installed to prevent shock damage.

Scale shall carry a minimum of a Five (5) Year Factory Warranty.    “Limited” Warranties shall be considered unacceptable.

Full scale accuracy shall be better than 1%.   Scale shall be CARBOY-SCALE™ with TUF-COATTM coating and

CENTURY® dial indicator, Model 4D____ as manufactured by FORCE FLOW, 2430 Stanwell Drive., Concord, CA 94520 USA (www.forceflow.com), or equal.

OPTION 1:	Scale shall have a SATELLITE® loop powered 4-20mA output proportional to gross weight.

OPTION 2:	Scale shall have a 5amp low level contact with manual reset button.

OPTION 3:	Scale shall have a 10amp fixed alarm pressure switch.

Model Numbers:

PLATFORM & INDICATOR:	Model No. (Capacity lbs)	Metric Model (Capacity kgs)		For Tank Diameter

4D200	(200 lbs)		4D100K	 (100 kg)		8" (20cm) - 16" (40cm)

4D400	(400 lbs)		4D200K	 (200 kg)		14" (35cm) - 19" (48cm)

4D600	(600 lbs)		4D300K	 (300 kg)		18" (45cm) - 24" (60cm)

File:   _WEB\SPECS\SPCCARHY   02/03/03

FORCE FLOW   1-800-893-6723

FAX: 925-686-6713

CARBOY/HYD