TYPICAL SPECIFICATION

SOLO© XT

55-GALLON DRUM SCALE

SPECIFICATION SECTION 11245 - Chemical Feed

SPECIFICATION SECTION 11240 - Chemical Metering Pumps

A quantity of ______ drum scales shall be provided and shall be of the digital readout/hydraulic load cell type.  For redundancy, each weighing platform shall have a sole and separate weight indicator with no shared components.  Scale shall be sized to accept drums or carboys from 18" (457mm) to 24" (610mm) in diameter.   Platform height shall be no more than 2" (51mm) to allow for easy handling and unloading of drums.  Platform scale coating system shall be a minimum dry thickness of 80 mils and be resistant to moisture, chemicals, abrasion, impact and UV light.

Scale shall be of the single hydraulic load cell design.  Load cell shall be of the temperature stable, rolling diaphragm type.  Flexible tubing shall connect load cell to indicator to allow easy remote installation of the digital indicator.  Tubing length shall be ______ ft. (meters) (6'/2M standard) in length.

Indicator shall utilize CROSS TECHNOLOGY to convert the hydraulic load cell signal into a digitally displayed weight value.  Indicator shall be battery operated and shall not rely on any type of external power for display operation.  Batteries shall provide approximately 4000 hours of display time.  Indicator shall be housed in a NEMA 4X enclosure with 6 function keys for indicator operation.  Net weight shall be displayed in two ways:  A numerical display and .81 inch high characters, and an analog 0-100% bar graph display.  A third display line shall prompt the user through routine user operations.

User shall be able to set the net weight either by scrolling in the net chemical weight, or scrolling in the tank tare weight.  A menu key shall provide access to the following 5 functions:   1) Zero Indicator, 2) Set bar graph & 4-20mA capacity, 3) Set display auto-off time, 4) Set alarm relay values and 5) Set decimal point.

OPTIONAL 4-20mA ANALOG OUTPUT:

Indicator shall have a Loop-powered (12-36 Volts DC by others) 4-20mA output proportional

to Net Weight.  An 18" flying lead shall be provided for termination in a user supplied

junction box.

OPTIONAL LEVEL ALARM RELAY:

Indicator shall have a .5 amp level alarm set point that can be configured for either a high

or low level alarm condition.  An 18" flying lead shall be provided for termination in a user

supplied junction box.

Scale shall carry a Full Five (5) Year Warranty.  “Limited” Warranties shall be considered unacceptable.

Full Scale accuracy shall be better than ½ of 1%.   Scale shall be DRUMM-SCALE™ with TUF-COAT™ coating and  SOLO© XT digital display, Model XT___DS___  as manufactured by FORCE FLOW, 2430 Stanwell Drive, Concord, CA 94520 USA (www.forceflow.com) or equal.

PLATFORM & INDICATOR:

Platform Material:			TUF-COAT				STAINLESS STEEL

Pounds			Metric		Pounds		Metric

BATTERY POWERED		XT10DS			XT5KDS		XT10DSS	XT5KDSS

WITH 4-20mA:			XT10DSMA		XT5KDSMA	XT10DSSMA 	XT5KDSSMA

WITH 4-20mA & RELAY:		XT10DSMAR 		XT5KDSMAR	XT10DSSMAR 	XT5KDSSMAR

File: _WEB\SPECS\SPC55GXT   02/03/03

FORCE FLOW   1-800-893-6723

FAX: 925-686-6713

55-Gallon DRUM/XT