TYPICAL SPECIFICATION

HYDRAULIC CENTURY®

AMMONIA CYLINDER SCALE

SPEC SECTION 11261 -   Ammonia Feed Systems

Cylinder scale shall be of the hydraulic cell type.  Scale shall be constructed of non-corrosive 1" (25mm) PVC plastic and sized to accept cylinders from 14" - 19" (356mm - 483mm) in diameter.  Platform height shall be no more than 1-5/8" (38mm) to allow easy handling and unloading of cylinders.     Platform scale coating system shall be a minimum dry thickness of 80 mils and be resistant to moisture, chemicals, abrasion, impact and UV light.

Scale cell shall be of the temperature-stable, rolling diaphragm type.    Flexible PVC coated copper tubing shall lead from the cell to allow easy wall mounting of the dial and shall be ____ ft. (meters) in length (6'/2M standard length).   Cylinder chaining bracket with integral tool rack shall be wall mounted and use a double coil chain and a spring loaded snap hook to secure cylinder.

Dial diameter shall be 4-1/2" (114mm) and read 0-400 lbs. (0-200 kg) net with provision for tare adjustment.  Dial shall be temperature stable with damper installed to prevent shock damage.Load cell system shall require no electric power and shall be immune to all RFI/EFI, power failures and lightening strikes.

Scale shall carry a minimum of a Five (5) Year Factory Warranty.    “Limited” Warranties shall be considered unacceptable.

Full scale accuracy shall be better than 1%.   Scale shall be an AMMONIA-SCALETM  400 with TUF-COATTM coating and CENTURY® dial indicator,  Model 4D_______-___AS  as manufactured by FORCE FLOW, 2430 Stanwell Drive, Concord, CA 94520 USA (www.forceflow.com), or equal.

OPTION 1:	Scale shall have a SATELLITE®  loop powered 4-20mA output proportional to gross weight.

OPTION 2:	Scale shall have a 5amp low level contact with manual reset button.

Model Numbers:

PLATFORM & INDICATOR:

No. of Cylinders		Model (Capacity lbs)		Metric Model (Capacity kgs)

1-Cylinder 		4D400-1AS  (400 lb)		4D200K-1AS  (200 kg)

2-Cylinder 		4D400-2AS  (400 lb/ea)		4D200K-2AS  (200 kg/ea)

File:_WEB\SPECS\SPCAMMHY   02/03/03

FORCE FLOW  1-800-893-6723

FAX: 925-686-6713    www.forceflow.com    info@forceflow.com

AMMONIA/HYD