TYPICAL SPECIFICATION

SOLO© XT

CARBOY SCALE

SPECIFICATION SECTION 11245 - Chemical Feeders

SPECIFICATION SECTION 11240 - Chemical Metering Pumps

A quantity of _____ carboy scales of _______ lbs. capacity shall be provided and shall be of the digital readout/hydraulic load cell type.    For redundancy, each weighing platform shall have a sole and separate weight indicator with no shared components.   Scale shall be sized to accept a ______ to _____ inch (cm) diameter carboy (see chart below).   Platform height shall be no more than 1-5/8" (38mm) to allow easy handling and unloading of carboys/drums.  Platform shall be constructed of 1" (25mm) PVC Plastic.   Platform scale coating system shall be a minimum dry thickness of 80 mils and be resistant to moisture, chemicals, abrasion, impact and UV light.

Scale shall be of the single hydraulic load cell design.  Load cell shall be of the temperature stable, rolling diaphragm type.  Flexible PVC coated copper tubing shall connect load  cell to indicator to allow easy remote installation of the digital indicator.   Tubing length shall be _____ ft. (meters) (6'/2M standard) in length.

Indicator shall utilize CROSS TECHNOLOGY to convert the hydraulic load cell signal into a digitally displayed weight value.  Indicator shall be battery operated and shall not rely on any type of external power for display operation.  Batteries shall provide approximately 4000 hours of display time.  Indicator shall be housed in a NEMA 4X enclosure with 6 function keys for indicator operation.  Net weight shall be displayed in two ways; a numeric display with .81 inch high characters, and an analog 0-100% bar graph display.  A third display line shall prompt the user through routine user operations.

User shall be able to set the net weight either by scrolling in the net chemical weight, or scrolling in the tank tare weight.   A menu key shall provide access to the following 5 functions: 1) Zero Indicator;   2) Set bar graph & 4-20mA capacity, 3) Set display auto-off time, 4) Set alarm relay values and 5) Set decimal point.

OPTIONAL 4-20mA ANALOG OUTPUT:

Indicator shall have a Loop-powered (12-36 volts DC by others) 4-20mA output proportional to Net

Weight.   An 18 inch flying lead shall be provided for termination in a user supplied junction box.

OPTIONAL LEVEL ALARM RELAY:

Indicator shall have a .5 amp level alarm set point that can be configured for either a high or low level alarm condition.  An 18 inch flying lead shall be provided for termination in a user supplied junction box.

Scale shall carry a Full Five (5) Year Warranty.   “Limited” Warranties shall be considered unacceptable.

Full scale accuracy shall be better than 1%.    Scale shall be CARBOY-SCALE with TUF-COATTM coating and SOLO© XT digital display Model XT_________ system, as manufactured by FORCE FLOW, 2430 Stanwell Drive, Concord, CA  94520 USA (www.forceflow.com), or equal.

PLATFORM & INDICATOR:		Model No	Metric Model	Scale		For Tank Diameter

BATTERY POWERED:	XT200		XT100K		Carboy 200	8" (20cm) - 16" (40cm)

XT400		XT200K		Carboy 400	14" (35cm) - 19" (48cm)

XT600		XT300K		Carboy 600	18" (45cm) - 24" (60cm)

WITH 4-20mA:       	XT200MA	XT100KMA	Carboy 200 	8" (20cm) - 16" (40cm)

XT400MA	XT200KMA	Carboy 400	14" (35cm) - 19" (48cm)

XT600MA	XT300KMA	Carboy 600	18" (45cm) - 24" (60cm)

WITH 4-20mA & RELAY:	XT200MAR	XT100KMAR	Carboy 200	8" (20cm) - 16" (40cm)

XT400MAR	XT200KMAR	Carboy 400	14" (35cm) - 19" (48cm)

XT600MAR	XT300KMAR	Carboy 600	18" (45cm) - 24" (60cm)

File:_WEB\SPECS\SPCARBXT    02/02/03

FORCE FLOW 1-800-893-6723

FAX: 925-686-6713

CARB/XT