TYPICAL SPECIFICATION

SOLO® G2 INDICATOR with

TANK WEIGHING ASSEMBLIES

SPECIFICATION SECTION 11245 - Chemical Feeders,

SPECIFICATION SECTION 11261 - Chlorination Equipment, or

SPECIFICATION SECTION 11240 - Chemical Metering Pumps:

A quantity of _____(1, 2, 3, or 4) tank weighing assemblies shall be installed for each tank/hopper/bin.  Each tank weighing assembly shall have a capacity of _______ lbs (kgs) for a total system capacity of _______ lbs (kgs).  Each tank weighing assembly shall consist of a nickel plated load cell of the shear beam strain gauge type and appropriate mounting hardware.  Flexible cable shall connect load cell to indicator for easy remote installation of the readout.  Cable length shall be _____ feet, 20' (6M) standard.

Tank weighing assemblies shall be self-checking and have provisions for thermal expansion and contraction.  Tank weighing assembly load plate shall be able to compensate for mounting surfaces up to 4 degrees off-level.

Indicator shall monitor__________(1 or 2 ) channels. The remote mounted LCD indicator shall carry CE marking and shall be housed in a NEMA 4X, UL approved enclosure. All operations shall be keypad operated & menu driven in order to avoid compromising the NEMA 4X seal at anytime.  The alphanumeric LCD readout shall have backlighting for readability in low light conditions.  Power requirement shall be ______________( 110 VAC, 220 VAC or 24 VDC)

A 6 digit numerical display shall give operator the ability to monitor chemical by weight (lb or kg) or volume (gallons or liters).  A bar graph display shall read 0-100% for the net contents.  A dual mode TARE key shall allow user to enter the tare weight of the vessel or enter the net weight of the chemical depending on application needs. A diagnostics menu shall allow recalibration without the need to apply field test weights.  A user adjustable filter function shall stabilize display in the event of vibration from pumps or mixers in the immediate vicinity of the scale.

Indicator shall output net weight via a 4-20mA signal and full scale output shall be user adjustable via the keypad.  Indicator shall have four adjustable set points to display low or high level conditions on the indicator.

OPTIONAL:    Set points shall activate ______ (up to four) ½ Amp dry contact relays.

Scale shall carry a Full Five (5) Year Factory Warranty.  “Limited” Warranties shall be considered unacceptable.   Pivoted weighing systems shall have an accuracy of 1/4 of 1%; all others 1/10 of 1%.     Weighing assembly shall be PROCELL®  Model PC-______ and digital indicator shall be SOLO® G2 Model SRG2-___ (1 or 2), manufactured by FORCE FLOW, 2430 Stanwell Drive, Concord, CA 94520 USA (www.forceflow.com).

Model Numbers:

INDICATOR:	SRG2-1		One Channel

SRG2-2 		Two Channel

MODEL NO	CAPACITY

TANK WEIGHING ASSEMBLY:

PC820		500 LB -10,000 LB (227 KG – 4563 KG)

PC160		1,000 LB. – 75,000 LB.  (454 KG – 34020 KG)

FORCE FLOW   1-800-893-6723

FAX: 925-686-6713    www.forceflow.com    info@forceflow.com

TWA/SOLOG2

02/26/09