TYPICAL SPECIFICATION

FOR BATTERY POWERED SOLO® XT WEIGHT INDICATOR

A quantity of _____ scales (cylinder, carboy or drum) shall be provided and shall be of the digital readout/hydraulic load cell type.  For redundancy, each weighing platform shall have a sole and separate weight indicator with no shared internal components.  Scale platform coating system shall be a minimum dry thickness of 80 mils and be resistant to moisture, chemicals, abrasions, impact and UV light.  Scale shall be of the single hydraulic load cell design.  Load cell shall be of the temperature stable, rolling diaphragm type.  Flexible tubing shall connect load cell to indicator to allow easy remote installation of the digital indicator.   Tubing length shall be _____ ft. (meters) (6' carboy/cylinder or 5’ drum standard) in length.

Indicator shall utilize CROSS TECHNOLOGY to convert the hydraulic load cell signal into a digitally displayed weight value.  Indicator shall be battery operated and shall not rely on any type of external power for display operation.  Batteries shall provide approximately 4000 hours of display time.  Indicator shall be housed in a NEMA 4X enclosure with 6 function keys for indicator operation.  Net weight shall be displayed in two ways; a numerical display with .81 inch high characters, and an analog 0-100% bar graph display.  A third display line shall prompt the user through routine user operations.

User shall be able to set the net weight either by scrolling in the net chemical weight, or scrolling in the tank tare weight.   A menu key shall provide access to the following 5 functions: 1) Zero Indicator, 2) Set bar graph & 4-20 ma capacity, 3) Set display auto-off time, 4) Set alarm relay values and 5) Set decimal point.

OPTIONAL 4-20 MA ANALOG OUTPUT:

Indicator shall have a Loop-powered (12-36 volts DC by others) 4-20 ma output proportional to Net Weight.  An 18 inch flying lead shall be provided for termination in a user supplied junction box.

OPTIONAL LEVEL ALARM RELAY:

Indicator shall have a .5 amp level alarm set point that can be configured for either a high or low level alarm condition.   An 18 inch flying lead shall be provided for termination in a user supplied junction box.

Scale shall carry a Full Five (5) Year Warranty.   “Limited” Warranties shall be considered unacceptable.

Full scale accuracy shall be better than 1%.    Scale Platform shall be TUF-COAT™ coating system protected and SOLO®  XT digital display indicator, Model XT_________  as manufactured by FORCE FLOW, 2430 Stanwell Drive, CA  94520 USA (www.forceflow.com), or equal.

PLATFORM & INDICATOR:		Model No		Metric Model	Scales		For:

BATTERY POWERED	XT150-1		XT100K-1	Cylinder			Single Cylinder Scale

XT150-2		XT100K-2	Cylinder		Dual Cylinder Scale

XT200		XT100K		Carboy 200	8” (20cm) – 16” (40cm)

XT400		XT200K		Carboy 400	14” (35cm) – 19” (48cm)

XT600		XT300K		Carboy 600	18” (45cm) – 24” (60cm)

XT10DS		XT5KDS		Drumm-Scale	30- to 55-Gal. Drums

WITH 4-20 ma:       	XT150-1MA	XT100K-1MA	Cylinder			Single Cylinder Scale

XT150-2MA	XT100K-2MA	Cylinder		Dual Cylinder Scale

XT200MA	XT100KMA	Carboy 200	8” (20cm) – 16” (40cm)

XT400MA	XT200KMA	Carboy 400	14” (35cm) – 19” (48cm)

XT600MA	XT300KMA	Carboy 600	18” (45cm) – 24” (60cm)

XT10DSMA	XT5KDSMA	Drumm-Scale	30- to 55-Gal. Drums

WITH 4-20 & RELAY:	XT150-1MAR	XT100K-1MAR	Cylinder			Single Cylinder Scale

XT150-2MAR	XT100K-2MAR	Cylinder		Dual Cylinder Scale

XT200MAR	XT100KMAR	Carboy 200	8” (20cm) – 16” (40cm)

XT400MAR	XT200KMAR	Carboy 400	14” (35cm) – 19” (48cm)

XT600MAR	XT300KMAR	Carboy 600	18” (45cm) – 24” (60cm)

XT10DSMAR	XT5KDSMAR	Drumm-Scale	30- to 55-Gal. Drums

File:_WEB\SPECS\SPCXTIND    02/18/03

FORCE FLOW   1-800-893-6723

FAX: 925-686-6713

XT/INDICATOR